Homework #5: GC Basics 25  pts total

Chem 6614 Chemical Instrumentation

Due Monday 30 March

Your name: _______answers_____________________________________

6.1. Terminology

- Name of common plumbing connection found in GC and HPLC _____SWAGELOK____

- The cone shaped item which is part of the above is called a ________FERRULE_______

3)  	Acronym for hypothetical distance for one fundamental step of separation_PLATE OR HETP

4)         Acronym for nondestructive `thermometer’ –like GC detector ____TCD____________

5)         Acronym for detector that burns sample, hydrogen and air _____FID________________

6)	What pushes the solvent in HPLC ?SINGLE PISTON `RECIPROCATING’ PUMP_______

7)	non-polar stationary phase/polar mobile phase=> _____REVERSED PHASE__________

8)	What you lay the stationary phase on (general name):_____SUPPORT___________

9)	Means constant pressure separation: ________ISOBARIC_____________________

10)      Chief measure of whether the peaks are apart: ___RESOLUTION___________

6.2. Technique matching

- You want to see the xylene (aromatic) contamination in a sample of methanol. Which GC detector is the best choice ? __PID  or ECD____________

- You have a mixture of complex amines dissolved in methanol which decompose easily when heated. What method would you use to separate the mixture ?___HPLC__

- A mixture of 4 alcohols with 5 carbons have boiling points ranging from 50-110oC. Which GC detector would you select to analyze them with ?__FID____________________

- The above mix produces only a broad and unresolved blob of a peak under isothermal conditions (T=70o )   . What could you do to improve resolution  and separation ?

THERMAL RAMP________________

- You want to analyze for the CO2 concentration in a gas sample. Which GC detector would

work best ? _____TCD_________________

f)	column type that can get up to 100 meters long; fragile; lowers required carrier gas pressure: __CAPILLARY________

6.3. Basic GC Calculations

- It takes a  flowmeter 30 seconds to pass 10 mL of He for a GC with a 6 foot, packed column.  What is the He flow rate in mL/minute ?  __________20 mL/min________

30 s = 0.5 min=>	10 mL/(0.5 min) = 20 mL/min

A symmetric peak appears at 12.0 cm with a width (~w0.1)  of 0.3 cm given a chart speed of 10 cm/minute for the above system.

- Compute the retention time in seconds:	 		____72___  = tR (sec)

tr = 12 cm/(10 cm min-1) = 1.2 min => 1.2 min*60 s/min=72 s

- Compute the # of plates for the column: 			_25,600 or 29,600___  = N

N= 16 or 18.6 * (tr/w)2  ; w = 0.3 cm/(10 cm min-1) * 60 s/min = 1.8 s

N = 16 or 18.5*(72/1.8)2 = 25,600 or 29,600

- Compute the height equivalent of the theoretical plate:_0.00715 cm or 0.00618 cm = H(cm)

H= L(col,cm)/N

L = 6 feet * 12 inches * 2.54 cm/in ~ 183 cm

H = 183/(25,600 or 29,600) =0.00715 cm or 0.00618 cm

6.4. Semi-Quantitative Analysis of a Mixture  (5 pts)

An initial chromatogram of a known mixture composed of 2% A, 2% B and 2% C (v/v) is recorded using the `wet’ needle sampling method. The retention times tr(k)   and areas, Ak,  in pA*s  for all three components (k=A,B,C) are below:

Observed Chromatographic results for Known Mixture of A,B and C (2% each v/v)

Answers  6.4a  sensitivities

Component		tr(min)		Ak		Sk = Ak/2%

A			0.456		95.43		47.72

B			0.654		187.39	            93.70

C			1.380		269.51	            134.3

An unknown mixture of the above components is next run under the same conditions as in the above known mixture. The following peak retentions and areas are observed

Observed Chromatographic results for Unknown Mixture of A,B and C

Answers 6.4b apparent V% = Ak /Sk

Component		tr(min)  	A

A			0.456		27.56		27.56/47.72=0.59 %

B			0.660		451.32		451.32/93.70=4.82 %

C			1.380		237.9		237.51/134.3=1.77 %

- Compute the apparent sensitivity of the known components S= (pA*s/%V)

- Compute the apparent volume % of A,B and C in the unknown mixture

- Assuming just three components appear in the chromatogram, predict the relative volume % of each component in the unknown mixture.

Answers to 6.4c

Relative % = apparent %  *100

 apparent %

Component		apparent %	Relative volume % calculated

A			0.59 %		0.59 *100/7.18=8.21

B			4.82 %		4.82*100/7.18 =67.1

C			1.77 %		1.77*100/7.18 =24.7

 apparent %		7.18

- Given that actual relative % volumes of the unknown mixture are as below, what is the average % error of the analysis over all three peaks ?

Actual Composition of Unknown mixture (relative Volume %)    calculated    %error1

A			  6.2        					   8.2		32.2

B		     	62.5       					67.1		  7.4

C			31.3      					24.7		-21.1

100 %							error%=-18.5

(don’t use absolute %)

1  (calc – actual) *100

Actual

Average error % = error%=-18.5/3 =6.2