Laboratory #6 (a `dry’ lab)

see also lab 6 at website

HPLC Analysis of aromatic mixture and estimation of Instrument Detection limit

Chem 6614	Instrumental Methods of Chemistry

SUNY Alfred State College

Due 29 March

HPLC Dry Lab Scenario: Part 1-Semi-quantitative mixture analysis  10 points

A standard 1:1:1 volume mixture of benzene, toluene and ethyl benzene was prepared by delivering 5 mL of each component into a 100 mL volumetric flask and diluting to the mark with methanol to make the standard PE 5% Reference mixture

The standard mixture was then loaded into the 20 uL injector port of the Perkin-Elmer (PE) Nelson 410 LC  (serial # 41 BN 2121405) equipped with a model 295 PE UV-VIS detector. A C-18 reverse phase, 2 inch, guard column was used to effect separation in a 30:70 mixture of water:methanol carrier solvent. The carrier flow rate was set a 1 mL/min and the analyzing wavelength was 260 nm using the FAST acquire spectrometer setting with A(min)=0.002.  Average pump pressure is ~900 psi.

The following retention times and peak areas (Abs*s units) are observed:

Table 1: Retention Times (tr) and HPLC  Peak Areas (A) Observed for 5% v/v Reference M ixture of Benzene/Toluene/Ethyl Benzene  (1:1:1) Standard Using Conditions Described

tr(min)		A(Abs*s)

0.636		626068

0.807		856639

0.997		812537

An unknown extraction containing some or all of the components in the standard is obtained next and a 20 uL sample of this unknown is loaded and run under the same conditions as applied to the 5 % solution of Benzene/Toluene/Ethyl Benzene  (1:1:1) Standard.

The following retention times and peak areas (Abs*s units) are observed:

Table 1: Retention Times (tr) and HPLC  Peak Areas (A) Observed for Unknown Extract in Methanol Obtained Using Conditions Described

tr(min)		A(Abs*s)

0.641		30021

0.800		76550

1.004		456032

Use the data provided above to estimate:

1)The absolute uL of each of the above components in the unknown 20 L sample used (Vk)

2) The relative volume % of each in the 20 uL sample injected. (%Vk)

Dry Lab Scenario: Part 2- Sensitivity Check

A series of 4 standard  dilutions (Dilute Mixtures #1-4)  of ethyl benzene have been prepared in methanol with the % volumes listed in Table 3. The observed area A (Abs*s) of their HPLC peaks at tr~ 1 minute are recorded using the conditions described in Part 1.  Note that the lowest % volume has been measured 3 times so that a standard deviation can be computed.

Table 3: Observed HPLC Peak Areas, A,  for Various  Dilute Concentrations of  Ethyl Benzene in

Methanol Using Conditions Described in Part 1 of Scenario

| Dilute Mixture # | % ethyl benzene (v/v) | Aobs(Abs*s) |
| --- | --- | --- |
| 1 | 0.050 | 8610 |
| 1 | 0.050 | 8334 |
| 1 | 0.050 | 8590 |
| 2 | 0.100 | 17001 |
| 3 | 0.200 | 33987 |
| 4 | 0.300 | 50980 |

Use the data to estimate the absolute detection limit, DL(ng), of the HPLC to ethyl benzene under the conditions applied.

The documentation in Experiment #6 to should be used to guide you in carrying out the analysis for both Part 1 and Part 2 of this scenario. Note that you are to hand in just the RESULTS section of what would have been Lab #6. (See page 6 of Lab #6).

Chem 6614  Instrumental Methods  of Analysis   HPLC Technique Questionnaire  10 pts

Your name: _________________________________________________

- What pushes the carrier in an HPLC ? _____________________________

- Why is the UV range of the detector most useful in HPLC analysis of organics ?

_____________________________________________________________________

- A fixed solvent mixture HPLC run is referred to as a(n) __________________ run.

- Why is a blunt syringe employed when loading a sample into the HPLC ?

____________________________________________________________________

- What device is used to remove the sinusoidal pressure variations intrinsic to HPLC ?

___________________

- What is being referred to in the HPLC column specification “300 ”________________

- What kind of carrier solvent is appropriate in a normal phase HPLC run ?______________

- Name two ways to improve (or at least change) the `SSS’ (Sharp-Symmetric-Separated)

character of an HPLC chromatogram.

1)____________________________________________________

2)____________________________________________________

- Why is HPLC more prevalent than GC in pharmaceutical labs for separating and analyzing mixtures ? _______________________________________________________

________________________________________________________________