## Slide 1: Tour of the Modern Magnetic Sector Instrument
Sector MS-University of Wisconsin Department of Chemistry
Inlet plumbing + source (North Carolina State instrument)
student

## Slide 2: Typical modern ms labs-au’ natural
U of Wisconsin
NC State
Johns Hopkins
Average ms scientist after a few years
MS…an extremely hardware-bound and often draining sport

## Slide 3
Average Mass Spectroscopist after a few years….

## Slide 4: Tour of common major ms components: magnets and “Conflat” UHV* Stainless steel plumbing
*UHV=Ultra High Vacuum
Close-up of magnet
Typical plumbing geography

## Slide 5
Conflat flange design

## Slide 6: Tour of major ms components: vacuum pumps
Roughing (mechanical) pump
Rough pump moves pressure from
1 atm- 0.00001 atm ~ 10 
Schematic of rough pump rotor assembly
Oil in here
Gas out here
Note on units:
1 mm Hg =1000 
1 mtorr= 1 
10-6 atm =0.76 mtorr ~ 1mtorr =1 

## Slide 7
ASC MECHANICAL PUMP FOR BOB
to turbopump

## Slide 8: MS components continued: diffusion pumps
Out to rough pump
1000   0.001 
Specially made` diffusion’ pump oil (a stable silicone, e.g. Dow-Corning 701) is heated, misted and rained down through nozzles to entrain gas from ms manifold
From ms manifold

## Slide 9
ASC OIL DIFFUSION PUMP ON ELANOR
TO MECHANICAL PUMP

## Slide 10: MS components continued: alternative to diffusion ..turbomolecular pumps
Cutaway of turbo pump
From ms manifold
Louvered vanes scoop up gas and guide them out to rough pump
Turbo pumps are now the dominant choice…not greasy and no outgassing problems…but $$$
Shaft turns at 104 /s
Nearly supersonic

## Slide 11: ASC TURBOMOLECULAR PUMP AND PHILLIPS GAUGE ON BOB

## Slide 12: MS components continued: turbomolecular Pumping speed vs pressure
1 
1000 
Oil diffusion begins to lose it here
Start of UHV (ultra high vacuum range: 10-8 atm)
Turbos rated down to10-6 

## Slide 13: MS components continued: other pump types
Cryopump
Fill cooling loop with refrigerant (liq. N2)
To rough pump (takes out air)
Achieves ~ 0.01-0.001 
and freeze out condensable gases on outside of loops

## Slide 14: MS components continued: other pump types (continued)
Titanium sublimation pump
(also called a getter pump) …two basic parts:
1) Ti is sputtered from three hair pin loops of Ti
3) Achieves 0.001 u ?
2)Reactive residual gases not pumped away by rough pump are reacted on baffle surface
A) Ti sublimation source
B) Getter surface

## Slide 15: MS components continued: other pump types (continued)
Ion pumps
1)Best choice for UHV work:
Ion trap collector
(+)
4) 10-10  10-12 atm
(0.01 -0.0001 )
2)Ions formed from impact of electrons which are gettered
3)or collected at anode trap

## Slide 16: MS components continued: vacuum gauges
Thermocouple gauge
Thermocouple junction measures variation of temperature created by residual gases and converts to volts measured at (M)
Poor chemist’ universal P gauge… $ 50 (without electronic amplifier)
Produces constant T at reference side
Good from 1 atm-10 

## Slide 17: MS components continued: vacuum gauges
Phillips-Granville high vacuum gauge
1)Works by sending a current across space between anode wire and cathode loop
cathode
anode
…and at fixed (high) voltage measuring conductance
Conductance measured here
Measures down 10-6 u
(also called simply: ion gauge …$1000/tube)

## Slide 18: MS components continued: Electron impact (EI) source
Various designs from HP, Thermoelectron and Perkin-Elmer

## Slide 19: MS components continued: Electron impact (EI) source (cont.)
Magnets used to focus beam and inject into ion accelerator*
…otherwise ions will wander like sheep
SOURCE BLOCK

## Slide 20: ASC SOURCE AND REPELLER FOR ELANOR
Filament off

## Slide 21
ASC SOURCE AND REPELLER FOR ELANOR
Filament on

## Slide 22: MS components continued: Electron impact (EI) source (cont.)
Ions created and focused out of source block slit are accelerated through a series of +/- plates
Addendum:
`accelerator’ to produce fixed velocities….tied to net voltage V across +/- plates
This piece caled the ION or MASS ACCELERATOR or `ION OPTICS’

## Slide 23
From GC
Alternative to EI: Chemical Ionization (CI):
A `soft’ alternative to EI
CI attributes
High pressure of `reagent’ gas (CH4 common)*
Low accelerating voltages
M +CH4+
CH4+ replaces electrons
But many other small molecules are also used: NH3, CH3CN (acetonitrile), C2H6…

## Slide 24
The low energy and `softness’ of CH4+ means much less fragmentation of target molecules
e- accelerated
at 1400 V
Fragmentation pattern
C6H6
Target molecule
C6H5+ C6 H4+ C5 H4+ C5H3+ C4H4+ C4H3 C3H3 C3H2…
CH4+
C6H6
CH4+ accelerated
at 70 V
Target molecule
EI
CI
C6H6+ or C6H7+
Only P+ or PH+

## Slide 25
Target molecule
MW 267
EI vs CI spectrum of a complex compound

## Slide 26
The quadrupole mass selector (ASC)

## Slide 27
ONE of ASC’s GC-MS (ELANOR)
GC
Transfer (AUX) line in here
MS
Filament and repeller (on)

## Slide 28
Side view of ASC MS (Babette): All you can see `under the hood without breaking vacuum

## Slide 29: The MS Body parts game
First ion source
… a stick shape
1st detector:
Kodak 35 mm film records spectrum
Mechanical pump
Where is the `detector ?
Where is the vacuum pump?
Where is ion source?

## Slide 30: Find the source game
White ceramic back cap is a hint

## Slide 31: MS parts game (cont.)
Diffusion pumps
magnet
(probably) source (or TC gauge ?)
ion acceleration (optics)
1)Find pumps
2) Find magnet
3) Find ion optics (accelerator)
4) Find ion source
or

## Slide 32: MS components continued: dynode detector design
Original electron dynode multiplier design: emissive plates at angles
Compact version of dynode multiplier
Gain factor is 106 –1012

## Slide 33: A few more MS bits & pieces: other designs and approaches
Alternative to dynode detector FARADAY CUP
Faraday cage
(keeps all charges inside)
Can also be voltage amplifier (op amp)
Detector pod

## Slide 34: A few more MS bits & pieces (continued)
Alternative to magnetic sector: quadrupole design
Pros
Smaller
Cheaper
Stable
Rugged
Cons
Tuning is tricky
Usually poorer resolution (R)
Alfred MS is quadrupole
See text for description of control issues

## Slide 35: A few more MS bits & pieces (continued)
Even cheaper alternative:
Time-of-Flight (TOF)
Pros (vs quadrupole)
Even smaller than quadrupole (no magnets at all !)
Con (vs quadrupole)
Resolution only 1 amu (lousy)
~ 1 foot1
1 even smaller than original Aston MS

## Slide 36: Summary of bits and pieces
Pump types
*Mechanical 1 atm -10 
*Turbomolecular 10-0.0001 
Diffusion 1000-0.001 
Cryopumps 10-0.01 
Ti sublimation 10 -0.01 
Ion pump 1-0.00001 
Gauge types
*Thermocouple 1 atm-10 
Phillips ion gauge 10 u ---0.00001 
Pirani gauge 100-0.1 
Most common source
*EI = Electron Impact
Most common detector
*Dynode multiplier (Rheodyne)
*Red colored components in Alfred instruments

## Slide 37
Example of Computing R= MS Resolution
28.05 amu
27.95 amu
28.00 = Mav
(half-way between two peaks)
Determine minimum spacing where two peaks are still distinguishable :
R= Mav
M
M=28.05-7.95=0.10
=28
0.1
Mass (amu)
count
=280 Higher is better

## Slide 38: MS Designs compared
Resolution = R= M/M
Example
for M=32 amu M=0.1 amu=>
R
$
Size
use
Magnetic sector
Max 20,000
300-500 k$
Need whole room
research
Quadrupole
Max 2000
75-100 k$
1-2 bench tops
Forensics, analytical chem, environment, education
TOF
Max <500
20-30 k$
½ bench top
Industry, limited species , field
Design feature
R= 32 = 320
0.1

## Slide 39: A few more MS bits & pieces (continued)
Indirect gas inlet
Direct probe
Traditional MS inlet designs

## Slide 40: GC-MS inlet system
He
GC oven
Capillary column
Quad head
(w/ion optics)
MS chamber
Transfer line (heated)
Quad rods
EI
Sample M
M+ (*)
*
*
To detector
Sample injection here

## Slide 41
To call a service person to fix a mass spec costs ~$2000-$7000/day (depending on parts)
So…You break it, you fix it, maggots !