## Slide 1
Even with fancy, computer-interfaced 21st century chemical instrumentation…things get ugly fast

## Slide 2
1st Ugly Fact of 21st Century Instrumentation Science:
Calibration of all modern instrumentation rests on 19th century volumetric/analytical methods
… with these tools
..we calibrate this tool

## Slide 3
21st century piece
19th century piece
Error determined by std. dev. of repeated measurements
Error ?...a problem…

## Slide 4
2nd ugly fact of 21st century instrumentation science: knowing the uncertainty of concentration requires icky math
Example: Primary stock1 solution concentration Co
What is the expected uncertainty in the concentration Co of 1.000 g NaCl in 50 mL volumetric flask (=0.02 g/mL)?
Mass error m=  0.001 g
Volume error V  0.05 mL
To correctly estimate of uncertainty in Co: 
Co = [(C/m)2(m)2 + (Co/V)(V)2)1/2
1Primary means weighed solute into calibrated volume

## Slide 5
% Co error = 100* Co/Co = 100*2*10-5/0.02
Co = [(C/m)2(m)2 + (Co/V)(V)2)1/2
=[(1/V)2*0.0012 + (m/V2)2*0.052]1/2
=[(1/50)2*0.0012 + (1/502)2*0.052]1/2 =2*10-5
Mass error piece
Volume error piece
=0.1%
Mass error m=  0.001 g
Volume error V  0.05 mL

## Slide 6
Error contributions to primary C0 = mass/volume
concentration for 1 g/50 mL stock solution
∆m(g) ∆V(mL) V error % m error % total % C error
0.0010 0.05 <1 >99  0.100
0.0001 0.05 10 90  0.011
Primary standard error comes mostly from the scale’s error
Technology sucks

## Slide 7
What about error in secondary standards
(diluting primary using pipets to a larger volume ?)
Example:
5 mL of primary 0.02 g/mL stock diluted 100 mL=>
0.02000 g/mL  0.00002 g/mL
5 mL pipet  0.01 mL
100 mL volumetric flask  0.08 mL

0.001000 0.000002 g/mL
secondary standard
After icky math…
Pipet primary Co 100 mL volumetric
71 % 18 % 12%
Distribution of error in secondary standard
Volumetric ware is the limit in secondary and later dilutions
(~0.2% error)

## Slide 8
Primary C0 error arises from balance ±m
Secondary Cs error arises from initial C0 error and pipet ±V. For a simple dilution error % in Cs is typically:
±0.2-0.3%
~ ±0.1%
Further dilutions increase the % error by ±(0.1-0.3%) for each additional dilution step. (Larger pipets better.)

## Slide 9
The chief source of solution error …”back in the day…”
m ~  0.001 g (10 swings)
A Becker Chainomatic 2-pan swing balance circa 1960
1) Fine tune balance with `chain’
2)Measure over ~10 swings with needle ending at same deflection left and right

## Slide 10
The chief source of solution error …now
Mettler-Toledo AK 160 g
4 place top-loading, taring electronic balance
m ~  0.001 g
While 0.0001 g balances are available at ASC, they tend to drift and are unstable unless great care is taken to stabilize the base and exclude drafts

## Slide 11
Secondary* solution concentrations, Cs
*secondary=> solute delivered as a volume from a primary standard C0 and diluted
Primary C0
Secondary Cs
Class A volumetric pipet transfers V(mL) of C0

## Slide 12
Class A Pipet Tolerances
Volume, V(mL) | ±V(mL) | % error
0.5 | ± 0.006 | 1.2
1.0 | ± 0.006 | 0.6
2.0 | ± 0.006 | 0.3
3.0 | ± 0.010 | 0.3
4.0 | ± 0.010 | 0.3
5.0 | ± 0.010 | 0.2
10.0 | ± 0.020 | 0.2
15.0 | ± 0.030 | 0.2
20.0 | ± 0.030 | 0.2
25.0 | ± 0.030 | 0.1
50.0 | ± 0.050 | 0.1
100.0 | ± 0.080 | 0.08

## Slide 13
The bottom line, kids….you will still need to be competent with all all those icky 19th century math things…. e.g., calculate in moles, ppm, wt % and do dilutions