## Slide 1: Tour of the Modern Magnetic Sector Instrument
Sector MS-University of Wisconsin Department of Chemistry
Inlet plumbing + source (North Carolina State instrument)
student

## Slide 2: Typical modern ms labs-au’ natural
U of Wisconsin
NC State
Johns Hopkins
Average ms scientist after a few years
MS…an extremely hardware-bound sport

## Slide 3: Tour of common major ms components: magnets and “Conflat” UHV* Stainless steel plumbing
*UHV=Ultra High Vacuum
Close-up of magnet
Typical plumbing geography

## Slide 4: Tour of major ms components: vacuum pumps
Roughing (mechanical) pump
Rough pump moves pressure from
1 atm- 0.00001 atm ~ 10 
Schematic of rough pump rotor assembly
Oil in here
Gas out here
Note on units:
1 mm Hg =1000 
1 mtorr= 1 
10-6 atm =0.76 mtorr ~ 1mtorr =1 

## Slide 5: MS components continued: diffusion pumps
Out to rough pump
1000   0.001 
Specially made` diffusion’ pump oil (a stable silicone, e.g. Dow-Corning 701) is heated, misted and rained down through nozzles to entrain gas from ms manifold
From ms manifold

## Slide 6: MS components continued: alternative to diffusion ..turbomolecular pumps
Cutaway of turbo pump
From ms manifold
Louvered vanes scoop up gas and guide them out to rough pump
Turbo pumps are now the dominant choice…not greasy and no outgassing problems…but $$$
Shaft turns at 104 /s
Nearly supersonic

## Slide 7: MS components continued: turbomolecular Pumping speed vs pressure
1 
1000 
Oil diffusion begins to lose it here
Start of UHV (ultra high vacuum range: 10-8 atm)
Turbos rated down to10-6 

## Slide 8: MS components continued: other pump types
Cryopump
Fill cooling loop with refrigerant (liq. N2)
To rough pump (takes out air)
Achieves ~ 0.01-0.001 
and freeze out condensable gases on outside of loops

## Slide 9: MS components continued: other pump types (continued)
Titanium sublimation pump
(also called a getter pump) …two basic parts:
1) Ti is sputtered from three hair pin loops of Ti
3) Achieves 0.001 u ?
2)Reactive residual gases not pumped away by rough pump are reacted on baffle surface
A) Ti sublimation source
B) Getter surface

## Slide 10: MS components continued: other pump types (continued)
Ion pumps
1)Best choice for UHV work:
Ion trap collector
(+)
4) 10-10  10-12 atm
(0.01 -0.0001 )
2)Ions formed from impact of electrons which are gettered
3)or collected at anode trap

## Slide 11: MS components continued: vacuum gauges
Thermocouple gauge
Thermocouple junction measures variation of temperature created by residual gases and converts to volts measured at (M)
Poor chemist’ universal P gauge… $ 50 (without electronic amplifier)
Produces constant T at reference side
Good from 1 atm-10 u

## Slide 12: MS components continued: vacuum gauges
Phillips-Granville high vacuum gauge
1)Works by sending a current across space between anode wire and cathode loop
cathode
anode
…and at fixed (high) voltage measuring conductance
Conductance measured here
Measures down 10-6 u
(also called simply: ion gauge …$1000/tube)

## Slide 13: MS components continued: Electron impact (EI) source
Various designs from HP, Thermoelectron and Perkin-Elmer

## Slide 14: MS components continued: Electron impact (EI) source (cont.)
Magnets used to focus beam and inject into ion accelerator*
…otherwise ions will wander like sheep
SOURCE BLOCK

## Slide 15: MS components continued: Electron impact (EI) source (cont.)
Ions created and focused out of source block slit are accelerated through a series of +/- plates
Addendum:
`accelerator’ to produce fixed velocities….tied to net voltage V across +/- plates (~”variac control”)
This piece caled the ION or MASS ACCELERATOR or `ION OPTICS’

## Slide 16: The MS Body parts game
First ion source
… a stick shape
1st detector:
Kodak 35 mm film records spectrum
Mechanical pump
Where is the `detector ?
Where is the vacuum pump?
Where is ion source?

## Slide 17: Find the source game
White ceramic back cap is a hint

## Slide 18: MS parts game (cont.)
Diffusion pumps
magnet
(probably) source (or TC gauge ?)
ion acceleration (optics)
1)Find pumps
2) Find magnet
3) Find ion optics (accelerator)
4) Find ion source
or

## Slide 19: MS components continued: dynode detector design
Original electron dynode multiplier design: emissive plates at angles
Compact version of dynode multiplier
Gain factor is 106 –1012

## Slide 20: Summary of bits and pieces
Pump types
*Mechanical 1 atm -10 
*Turbomolecular 10-0.0001 
Diffusion 1000-0.001 
Cryopumps 10-0.01 
Ti sublimation 10 -0.01 
Ion pump 1-0.00001 
Gauge types
*Thermocouple 1 atm-10 
Phillips ion gauge 10 u ---0.00001 
Pirani gauge 100-0.1 
Most common source
*EI = Electron Impact
Most common detector
*Dynode multiplier (Rheodyne)
*Red colored components in Alfred instruments
(HP has diffusion pump;
INCOS has turbo pump)

## Slide 21: A few more MS bits & pieces: other designs and approaches
Alternative to dynode detector FARADAY CUP
Faraday cage
(keeps all charges inside)
Can also be voltage amplifier (op amp)
Detector pod

## Slide 22: A few more MS bits & pieces (continued)
Alternative to magnetic sector: quadrupole design
Pros
Smaller
Cheaper
Stable
Rugged
Cons
Tuning is tricky
Resolution 
Alfred MS is quadrupole
See text for description of control issues

## Slide 23: A few more MS bits & pieces (continued)
Even cheaper alternative:
Time-of-Flight (TOF)
Pros (vs quadrupole)
Even smaller than quadrupole (no magnets at all !)
Con (vs quadrupole)
Resolution only 1 amu (lousy)
~ 1 foot1
1 even smaller than original Aston MS

## Slide 24: A few more MS bits & pieces (continued)
Indirect gas inlet
Direct probe
Traditional MS inlet designs

## Slide 25: GC-MS inlet system
He
GC oven
Capillary column
Quad head
(w/ion optics)
MS chamber
Transfer line (heated)
Quad rods
EI
Sample M
M+ (*)
*
*
To detector
Sample injection here

## Slide 26: MS Designs compared
Resolution = R= M/M
Example
for M=32 amu M=0.1 amu=>
R
$
Size
use
Magnetic sector
Max 20,000
300-500 k$
Need whole room
research
Quadrupole
Max 2000
75-100 k$
1-2 bench tops
Forensics, analytical chem, environment, education
TOF
Max <500
20-30 k$
½ bench top
Industry, limited species
Design feature
R= 32 = 320
0.1