## Slide 1
Typical Corning Glass Batch Sheet for Experimental Glass Melt

## Slide 2
Glass Composition Phase Diagram
Underlying crystalline habits that characterize the various glass formation regions vs quench rate

## Slide 4
Effect of soft lithia glass in silica matrix after HF etch
X22,400 (SEM)
Examples of bi-phasic glass; example 1-Vycor

## Slide 5
Examples of bi-phasic glass; example 2
Pyrex (>50,000X)
(silver ions have been `interculated’ into structure to reveal biphasic character)
Pyrex is a combination of `softer’ glass (so it flows under flame )and a much harder glass (so it resists chemical and physical change)

## Slide 6: Scanning Electron Microscopy (SEM)
Optical Microscopy
(obs) ~ 500 nm ~minimum object size
1000X practical maximum magnification
A. Leeuwenhoek 1632-1723

## Slide 7
Electron microscopy: a breakthrough microscopy
=h/p (De Broglie equation) ~ 0.02 nm for 3.6 keV beam
Practically=>Can see 0.2-1 nm objects
higher beam energy, lower 
40-100,000 X magnifications
Manfred v. Ardenne 1907-1997
invents in 1937
Charles Oakley
1904-1996
develops in 1947

## Slide 8: BASICS of SEM technique
Main imaging methods
Secondary electrons (SE)
Backscattered electrons (BSE)
Main analytical method
XRED = EDX (same)
(X-Ray Energy Dispersion or Energy Dispersive analysis of X-rays
1-10 keV
Beam spot typically 0.2 nm)

## Slide 9: Some Geeky Details Of SEM
Inelastic BSE =impacting electrons re-scattered (sensitive to atomic masses producing scattering)
Auger (pronounced OJay)=surface state electrons
SE = electrons re-emitted by absorbing materials (provides great depth of field & primary image)
Characteristic x-rays= energy spacings of elements (used for XRED=EDX)
Electron beam produces a characteristic tear-drop shaped hole in material, which vomits electrons and x-rays
5-100 um
0.2-2 nm spot

## Slide 10: More geeky details of SEM:SE detection

## Slide 11: Typical `SE’ images
Cheese at 1500X
Spider mite
Hair on spider mite’s leg
Ballatidium coli (stomach protozoan) at ~5000X

## Slide 12: Yet more geeky details: BSE & EDX
Magnetic doughnut used to gather BSE and direct into SE (Everhart-Thornley) detector
BSE electrons
BSE provides map of element and compound distributions
(but poor depth of field)
BSE=backscattered electrons

## Slide 13
Basic SEM
XRED (x-ray energy dispersion unit)…an add-on unit
SEM and accompanying XRED spectrum of scanned region
Elemental analysis EDX
Sampling Limitation
Sample volume and area depend on beam
Top not like insides ?

## Slide 14: The last few details on SEM
Metals and other conductors can be observed without preparation
Non-conductors (glasses, ceramics, biological and organic samples) need to be coated with a conductive thin layer or the beam will cause local `melt’ down at beam focus and sample will glow
+ -
Au or C
Evacuated chamber
Filament power
Coated sample
Schematic of an SEM coater
Initially uncoated sample

## Slide 15: DSC (Differential Scanning Calorimetry)*
Schematic of basic DSC/DTA
Thermocouple pair (measures T difference)
Heated strip
(can go to 1800 C)
DSC if heat difference measured;
DTA (differential temperature analysis if temperature difference measured)
measures exo or endothermic (sample-reference) response vs input energy (or temperature) of strip
Power supply/interface

## Slide 16: Glass transition (Tg) and crystallization (Tc) (=devitrification)
Tg Tc
Glass melt is quenched faster than structure can reorganize into crystallographically ordered arrays=> quenched glass has no thermodynamic definition…glass is not a liquid (which is defined)
Disorder is a function of quench rate
Supercooled liquid is semi-ordered and thermodynamically defined
It takes energy to `reorganize’ chaos into supercooled order=> why Tg is endo proces
T
Q(cal)
‘supercooled liquid structure’
Disordered glass
Ordered crystal structure
Tg
Tc

## Slide 17: Kinds of thermograms
Single phase, melting
(e.g. cubic NaCl)
Crystal-crystal transition followed by melting (tridymite silica->cristobalite silica)
Release of hydration water (e.g CuSO4*5H2O)
Simple glass thermogram
(e.g. glassy silica->cristobalite)
mp (endo)
Phase change (exo)
Melting
Dehydration (exo)
melting

## Slide 18
Some useful facts to help sort out DSC curves
1) Typical melting points for selected materials
MATERIAL MELTING POINT, O C
Pyrex glass 1250
Soda lime glass 1000-1050
Most inorganic salts 600-900
(NaCl, CaCl2, KCl)
Plastics <400
Silica (glass or crystal) ~1800

## Slide 19
More useful facts to help sort out DSC curves
2) Glass transition temperatures (Tg) for selected glassy materials
Glass Tg O C
Silica (glass) ~1200
Pyrex glass 550
Soda-lime glass 500- 550
Plastics < 100

## Slide 20
And yet more useful facts to sort out DSC curves
1) Waters of hydration make multiple sharp, (weak) exothermic peaks
2) crystal-crystal transitions (Tcc) make pretty sharp, (strong) exothermic peaks
3) glass-crystal transitions (Tgc) make broad, (strong) exothermic peaks
4) For a glass, Tg < Tgc < Tcc

## Slide 21
Possible Materials Producing Exercise 9 DSC (modified)
PVC plastic
KCl (mp 740 oC)
NaCl (mp 820 oC)
CuSO4 *5H2O
Soda-Lime Glass
Pyrex glass
Silica glass

## Slide 22
Using the list on the previous slide, try to rationalize the composition of the samples measured in thermograms A-F
Exercise 9: Interpreting DSC Curves (modified)

## Slide 23: XRD (X-ray Diffraction)
Basic `X-ray powder diffraction’ camera (Debye-Scherrer design)
Bragg’s Law: 2*d*sin=
Given fixed  & a measure of the angle , you can determine atomic spacing d
The key point….
X-rays in
X-rays in
Diffracted X-rays
Diffraction pattern of DNA
Taken by Rosalind Franklin

## Slide 24: Details of Debye-Scherrer XRD diffraction camera
rings
or
lines

## Slide 25: Examples of XRD spectra
Pure crystalline substance
Glass only (`glass halo’)
Partially `devitrified’ glass*
Only lines…all crystal
Halo + lines= glass + crystal (devitrified glass)
Halo only= all glass